## Slide 1
10. Noise, Distortions, and the MNF Transform
Dr. David R. Thompson
Jet Propulsion Laboratory, Imaging Spectroscopy Group
July 20, 2019

## Slide 2
Goals for this module
Recognize different flavors of instrument noise and systematics
Estimate instrument noise from imaging spectroscopy data
Calculate the Minimum Noise Fraction (MNF) transformation

## Slide 3: Physical instruments produce various flavors of distortion:
Grating
Slit
Focal Plane
Array
Telescope
Photon shot noise: magnitude is square root of signal due to Poisson counting statistics
Electronic read noise: magnitude is usually independent of signal
Nonlinearities: departure of radiometric response from idealized proportional description
Other Systematics: Optical ghosting, electronic panel effects, calibration error, etc.
Highly instrument-specific.
Our emphasis

## Slide 4: Estimating Noise Part I: First Principles Method
Grating
Slit
Focal Plane
Array
Telescope
In this approach we calculate radiometric uncertainty by forward propagation of photon counting and electronic noise statistics
We fit models by analysis and/or laboratory characterization
Note that Signal to Noise Ratios (SNRs) are signal-dependent in shot noise-limited systems
Pro: SI-traceability; component-wise models; high accuracy
Con: In-flight validation & updates difficult, pathological illumination, tricky to get covariance structure

## Slide 5: Estimating Noise Part II: Scene-based noise estimation
This alternative estimates radiometric uncertainty directly using empirical statistics of the data.
It is easier than vicarious characterization, because it does not require in situ data. But it does typically require a spatially homogeneous target area so it is not suitable for all scenes.
We use Von Neuman’s approximation based on pairwise differences:
…applied separately in each channel, preferably along pushbroom (down-track) dimension to characterize each element independently.
[Von Neumann, J. Ann. Math. Statist. 12 367–395. (1941). MR0006656]

## Slide 6: Scene-based noise estimation (continued)
[From Thompson et al., RSE 2019]
Crudely speaking, the differences between neighboring locations in homogenous areas represent samples of the instrument noise.
This can be used to construct sample-based covariance matrices using the approaches presented previously.
The covariance matrix defines a zero-mean multivariate Gaussian noise distribution.

## Slide 7: Scene-based noise estimation (continued)
If the scene is not spatially homogeneous, it might still be locally smooth. This leads to an alternative based on nonparametric smoothing.
Estimate the ”scene contribution” by convolving a local smoothing operator (e.g. a Gaussian) and subtract that signal.
Then estimate noise with our empirical expression for the sample covariance matrix:
Original Image
Subset (550 nm)
Scene, smoothed in the downtrack dimension
“Noise contribution”
is the remainder
[Thompson et al., Optics Express 2017]
Total = Scene + Noise

## Slide 8: The MNF transformation
As detailed in: A. A. Green, et al, IEEE Trans. Geosci. Remote Sens., vol. 26, no. 1, pp. 65–74, Jan. 1988.
Make the signal zero mean (subtract µ from each spectrum)
Calculate the eigenvalues and eigenvectors of the noise covariance, perhaps with SVD
“Whiten” the noise to remove influence of noise covariance
Apply Principal Component Analysis
Result: PCA, but scaled so that noise level is unity
Noise-whitened data matrix
Diagonal matrix, reciprocal of square root of N’s eigenvalues
Eigenvectors of noise covariance N
Zero-meaned data matrix
Eigenvectors of noise-whitened covariance matrix
Data in MNF projection
The Minimum Noise Fraction (MNF) transformation is a linear orthogonal basis similar to PCA, but forcing noise in each channel to be unity. It is actually the product of two rotations, e.g. noise-”whitening” and PCA:

## Slide 9: Typical SNR curves
Image From Mouroulis et al., Optics Express, 2016
Signal to Noise Level (SNR) is one of several ways to describe noise.
The Noise-equivalent change in radiance (NeDL) is another.
Recall that both are dependent on the signal level – this is because photon shot noise dominates illumination-independent terms as more photons are counted.

## Slide 10: Tuned grating efficiencies
From Mouroulis et al., OE 2016

## Slide 11: Bonus: accounting for relative calibration drift
This panel shows in-flight estimates of calibration drift in the PRISM imaging spectrometer. Vertical lines relate to FPA electronic effects.
Estimating this from flight data permits updated synthetic flat field and per-element instrument noise estimation. We image a uniform target area, and then calculate the systematic differences between neighboring pixels on the focal plane array.
[Thompson et al., Remote Sensing of Environment 231, 2019]

## Slide 12: There are many other potential error sources
Raw Digital Numbers
Radiance at sensor
Surface Reflectance
Geophysical Map
Quantization error (bit rate)
(Influenced by observation noise)
Atmospheric unknowns (aerosol types, gas and particle vertical profiles)
Bi-directional reflectance
Topographic effects
Temporal variability
Retrieval models
Generalizing local parameterizations
Diverse in-situ data sources
Detector sensitivity drift
Transient or systematic electronic effects
Long term degradation and contamination
Spectral response uncertainty
Investigators should be aware of the uncertainties at each step in their analysis:

## Slide 13: Climatological uncertainties: AOD550
ENF
EBF
DF
DBF
MF
CS
OS
WS
S
G
PW
C
U
CM
B
Land cover classification
MODIS average AOD550
[Thompson et al., Remote Sensing of Environment 232, 2019]
Atmospheric correction may also introduce errors, particularly in conditions with high Aerosol Optical Depth (AOD).
AOD is the Beer-Lambert extinction coefficient due to aerosols, typically reported at a single wavelength like 550 nm.
Certain regions of the world have more atmospheric aerosols than others.

## Slide 14: Climatological uncertainties: water vapor
ENF
DF
DBF
MF
CS
OS
WS
S
G
PW
C
U
CM
B
EBF
MODIS average water vapor gm cm-2
Land cover classification
[Thompson et al., Remote Sensing of Environment 232, 2019]
Aerosols are particularly challenging when present with high water vapor content in the atmosphere.
The next section describes new atmospheric correction methods intended to handle these tricky combinations.

## Slide 15: Atmospheric correction artifacts
From Clark et al., 2002.
It is common for reflectance spectra to show spikey “residuals” in water vapor features where the model does not quite match true atmospheric conditions.
Look to the red arrows (940 nm and 1140 nm bands) for these issues.
The 1380 and 1880 nm dips are due to extreme absorption (the atmosphere is opaque). These regions are typically ignored.

## Slide 16: Atmospheric aerosol-related artifacts
Path radiance causes an upslope near 400 nm
Reduced contrast lowers height of 720 nm “red edge”
Aerosols distort the H2O vapor absorption shapes causing residual “spikes”
Different aerosol optical types have different effects on the spectrum. Some tend to absorb strongly (e.g. soot) while others are “scattering” (e.g. some dust). Their signatures may have different shapes depending on the specific particle shape, composition, and size.
Here we see some additional aerosol artifacts that appear in a vegetation reflectance profile.