## Slide 1
8. Radiances
Dr. David R. Thompson
Jet Propulsion Laboratory, Imaging Spectroscopy Group
July 20, 2019

## Slide 2
Goals for this module
Describe factors affecting the physical measurement of radiance spectra
Interpret and apply calibration data to uncalibrated spectral data cubes
Validate radiances in remote field data using in-situ measurements

## Slide 3: Reprise
AVIRIS-NG Installation

## Slide 4: Conventions
Samples (cross-track locations)
Rows (spectral channels)
Spectrometer
Telescope
FPA
Slit

## Slide 5: Working backwards from the FPA
Spectrometer
Telescope
FPA
Slit
Account for electronic effects - the unique, time-dependent radiometric response of each FPA element
Account for optical effects - the spatial and spectral “view” of each FPA element
Tie to the S.I. (absolute spectroradiometry)

## Slide 6: Working backwards from the FPA
Account for electronic effects - the unique, time-dependent radiometric response of each FPA element
Account for optical effects - the spatial and spectral “view” of each FPA element
Tie to the S.I. (absolute spectroradiometry)
Correct dark offset
Correct pedestal shift
Correct panel ghost
Correct flat field
Correct crosstrack scatter
Correct spatial scatter
Apply radiometric coefficients
Correct bad pixels
Statistical destriping

## Slide 7: Basic radiometric model
Radiance at FPA
uW nm-1 sr-1 cm-2
FPA Value
DN
Offset
DN
For each crosstrack element x and spectral channel c
Gain
uW nm-1 sr-1 cm-2 DN-1

## Slide 8: Basic radiometric model
Offset
DN
For each crosstrack element x and spectral channel c
Gain
uW nm-1 sr-1 cm-2 DN-1
Dark Level
(per element, each flightline)
Pedestal shift
(per spectrum, each frame)
Radiometric Coefficient
(per channel, each campaign)
Flat field
(per element, each campaign)

## Slide 9: Components of offset
Estimate dark level from initial non-illuminated segment of flightline
Estimate pedestal shift from non-illuminated channels

## Slide 10: Components of gain
Estimate RCCs using a NIST-calibrated source
Estimate flat field from a uniform sphere passed across the FOV

## Slide 11: Estimate RCCs using a NIST-calibrated source
Chapman et al., Remote Sensing (submitted)

## Slide 12: Estimate RCCs using a NIST-calibrated source
Chapman et al., Remote Sensing (submitted)

## Slide 13: Estimate flat field from a uniformly-illuminated spherical cavity passed across the Field of View (FOV)
Chapman et al., Remote Sensing (submitted)

## Slide 14: Validation method
RTM
images: Mishra et al., Heliyon 2015, wikimedia
reflection
scattering
absorption

## Slide 15: Validation
AVIRIS-NG remote measurement
Model from In-situ measurement
Residual
AVIRIS-NG
MODTRAN simulation

## Slide 16: Other artifacts
milliseconds
minutes
months
Whole FPA
FPA sub-region
FPA element
Flat field
NIST radiometry
Atmospheric wavelength tuning
Pedestal shift
Dark offset
Markov destriping
Panel Ghost
years
Laboratory wavelength calibration
SRF correction
“Bad element” replacement
Electronic and thermal effects
Detector and alignment
Optical effects

## Slide 17: A comprehensive correction strategy
Electronic effects - the time-dependent radiometric response of each detector
Optical effects - the spatial and spectral “view” of each detector
Calibration to the S.I. (absolute spectroradiometry)
Correct dark offset
Correct pedestal shift
Correct panel ghost
Correct flat field
Correct crosstrack scatter
Correct spatial scatter
Apply radiometric coefficients
Correct bad pixels
Statistical destriping

## Slide 19: Exercise 7: Radiance Calibration / Validation
Objective: Calibrate digital numbers and “close the loop” with predicted radiances from surface measurements acquired coincidently with an overflight of Ivanpah Playa, NM.. Compare alternative Radiative Transfer Models (RTMs). Which is best? Files include:
predicted_radiance_atrem - a prediction of at-sensor radiance based on the approach of Thompson et al., (RSE 2015).
predicted_radiance_modtran - via the MODTRAN Radiative Transfer Model.
flat - a flat field frame
rcc.txt - a file of radiometric calibration coefficients
dark - a dark current frame
dn - an imaging spectrometer file containing raw spectrometer data
loc – a three band matrix containing latitude, longitude, and altitude for the (non-orthorectified) dn file.
Reflectance was measured between two blue tarps placed on the playa surface, and used in the radiative transfer model. The locations:
South tarp - 35.60285N, -115.40388W
North tarp - 35.60466, -115.40266