## Slide 1
8. Spectral Calibration
Dr. David R. Thompson
Jet Propulsion Laboratory, Imaging Spectroscopy Group
July 20, 2019

## Slide 2
Goals for this module
Describe common distortions caused by physical misalignments in spectrometers
Recognize, characterize and correct spectral calibration error in flight data

## Slide 3: Conventions
Samples (cross-track locations)
Rows (spectral channels)
Spectrometer
Telescope
FPA
Slit

## Slide 4: Flavors of non-uniformity
courtesy
Robert O. Green, JPL
Wavelength
Cross Track Sample
Requirement
Failure by Frown
Failure by Twist
Failure by Spectral-IFOV-Shift
Wavelength

## Slide 5: Wavelength fitting: use the atmosphere as an in-flight calibration standard
Courtesy Robert O. Green

## Slide 6: Wavelength tuning with the atmosphere
Solar zenith
1. Calculate an average Top of Atmosphere reflectance for each cross-track position:
Measured radiance
Solar irradiance

## Slide 7: Wavelength tuning (contd.)
width
2. Calculate absorption coefficients using an appropriate Radiative
Transfer Model (I use 6s or MODTRAN).
3. Optimize continuum and absoprtion depth via Levenberg-Maquardt gradient descent:
shift
A linear continuum

## Slide 8: Wavelength fitting example
Apparent reflectance
Residual error
Apparent reflectance
Residual error
Common features used for this purpose include the oxygen A band (760 nm) and Near-Infrared H2O vapor absorptions
[Thompson et al., Atmos. Meas. Techniques 2015]

## Slide 9: Uniformity achieved
(courtesy Robert O. Green, JPL, and others)
Hyperion 2000
CRISM 2005
JPL M3 2008
AVIRIS-NG
[Thompson et al., AMT 2015]

## Slide 10: “Stray” spatial / spectral response functions

## Slide 11: Measurement model
(Lpred C)T S = Lmeas
[d x n]
[n x n]
[d x d]
[d x n]
Surface Reflectance
Atmospheric Attenuation, Scattering
“Extra” CRF
Predicted Radiance
At Focal Plane Array
Measured Radiance
At Focal Plane Array
Ideal PSF
Solar Irradiance
Scene model
“Extra” SRF
n = number of crosstrack samples
d = number of spectral channels
C = crosstrack stray light
S = spectral stray light

## Slide 12: Method
Find well-constrained properties of scenes’ true radiance
Posit the functional forms of the CRF and SRF
Optimize free parameters to match observations via:
Correct future data using the following linear transformation:
Here, + represents the Moore-Penrose inverse, e.g.
(Lpred C)T S = Lmeas
[d x n]
[n x n]
[d x d]
[d x n]
Lcorr = ((CT)+ ((ST)+ LmeasT ) T ) T
C+ = (CT C)-1 CT
C+C = I

## Slide 13: Procedure
Exploit the predictable shape of the O2 A band
Find a haze-free day to constrain path radiance
Calculate “true” A band based on elevation and sensor altitude
Dataset: Death Valley transect, a large elevation gradient
A band
3200 m
-80 m
N
4000 m
Surface Elevation

## Slide 14: Atmospheric fitting of PSF Tails
Spatial dimension
Spatial dimension
Spatial dimension
Spectral dimension
Spectral dimension
Spectral dimension
Nominal Response Function
In-Flight SRF Estimate via Atmospheric Fitting
Laboratory Measurement of Subpixel Stimulus
After SRF Correction
Pressure Altitude Retrieval Before SRF Correction
Results from Thompson et al., RSE 2018

## Slide 16: Exercise 8: Spectral Calibration
Objective: Use the atmosphere as a calibration object. Fit a modeled transmission spectrum to TOA observations.
Files include two examples of TOA reflectances acquired by the AVIRIS-NG instrument onboard a b-200 aircraft over Ivanplah playa. The exercise notebook contains fitting routines, and a transmission spectrum for appropriate observing geometry and water vapor loadings.
Pick an atmospheric feature, and eyeball an initial calibration. Start with a spectral response function that is 5-6 nm Full Width at Half Maximum. Then, fit wavelength shifts to fine tune your initial guess. Is there a way of quantifying how accurate this method can be?